| PRODUCTION (Electronics Systems) |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DETAIL PRODUCTION PLAN : MONTH / YEAR: |  |  |  |  |  |  |  |  |  |  |  |
| JOB ORDER NO: | CUSTOMER NAME | SYSTEM | QTY |  | OPERATIONS AND PLANNED DATE OF COMPLETION |  |  |  |  |  |  |
|  |  |  |  |  | MATERIAL | ASSEMBLY | WIRING | TESTING | LOAD CELL | FOR Q.C | REMARKS |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
|  |  |  |  | PLANNED |  |  |  |  |  |  |  |
|  |  |  |  | ACTUAL |  |  |  |  |  |  |  |
| APPROVED BY HOD PDN |  |  |  |  |  |  |  |  |  |  |  |

Form No.  : R/PP-05/01/REV 00 Pg 1 of 1

JOB TRAVELLER CARD – PRODUCTION  SYSTEMS

JOB No  :					               CUSTOMER:

SYSTEM:						   QTY:

| ASSEMBLY Doc. No Chassis Level NAME: Starting date & time: Completion date & time: signature: Final assembly NAME: Starting date & time : Completion date & time: Feed back to design, If Any: signature: Inspected by: | REWORK |  |
| --- | --- | --- |
|  | Reasons and Nature of Rework |  |
|  |  | Carried out By: Time Taken: Inspected by: |
| Wiring NAME: Doc. No: Starting date & time : Completion date & time: Feed back to design, If Any: Signature: Inspected by: |  | Carried out By: Time Taken: Inspected: |
| Testing NAME: Doc. No: Starting date & time : Completion date & time: Given to Q.C. on: Feed back to design, If Any: Signature: |  | Carried out By: Time Taken: Inspected: |

Note : In case of more than one rework use another form for the rework.

Form No.  : R/PP-05/03/REV 02 Pg 1 of 1

GENERAL ACCELEROMETER - JOB TRAVELLER

Model No…………………………          Sl. No. ……………………………..

Customer …………………………           Job. Order No……………....Supply Date……………...

Crystal Details   i) Batch………. (ii) Size – ID ……..mm, OD……….mm, Thickness……….mm

Weight of Accelerometer parts in gms. a) Mass………. b) Base………… c) Spring……………….

Total Weight………….. gms ,      Height…………….. mm , Spanner size ………………….mm

Resonant Frequency ………….Hz , Leakage Resistance…………MΩ , Capacitance…………..pF

| Sl. No. | Work Progress | Status | Work Done By | Sig. | Date | Work Insp. by | Sig. | Date |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | VISUAL INSPECTION - Crystal Seating - Spring Mechanical Play - Finish |  |  |  |  |  |  |  |
| 2 | Ultra – sonic Cleaning |  |  |  |  |  |  |  |
| 3 | Connector Fixing |  |  |  |  |  |  |  |
| 4 | Cleaning |  |  |  |  |  |  |  |
| 5 | Lacquer Coating on Base Shaft |  |  |  |  |  |  |  |
| 6 | Cleaning all parts for Assembly |  |  |  |  |  |  |  |
| 7 | Crystal Assembly |  |  |  |  |  |  |  |
| 8 | Preliminary Test |  |  |  |  |  |  |  |
| 9 | Spring Sealing |  |  |  |  |  |  |  |
| 10 | Calibration |  |  |  |  |  |  |  |
| 11 | Cover seating |  |  |  |  |  |  |  |
| 12 | Final Calibration |  |  |  |  |  |  |  |

Charge Sensitivity of the Accelerometer…………………………….pC/g

Verified By:				Approved By:			(HOD, PDN)

Date:

Form No.  : R/PP-05/05/REV 00 Pg 1 of 1

INDUSTRIAL ACCELEROMETER - JOB TRAVELLER

Model No…………………………          Sl. No. ……………………………..

Customer …………………………           Job. Order No……………....Supply Date……………...

Crystal Details   i) Batch………. (ii) Size – ID ……..mm, OD……….mm, Thickness……….mm

Weight of Accelerometer parts in gms. a) Mass………. b) Base………… c) Spring……………….

Total Weight………….. gms ,      Height…………….. mm , Spanner size ………………….mm

Resonant Frequency ………….Hz , Leakage Resistance…………MΩ , Capacitance…………..pF

| Sl. No. | Work Progress | Status | Work Done By | Sig. | Date | Work Insp. by | Sig. | Date |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | VISUAL INSPECTION - Crystal Seating - Spring Mechanical Play - Finish |  |  |  |  |  |  |  |
| 2 | Ultra – sonic Cleaning |  |  |  |  |  |  |  |
| 3 | Connector Fixing |  |  |  |  |  |  |  |
| 4 | Cleaning |  |  |  |  |  |  |  |
| 5 | Lacquer Coating on Base Shaft |  |  |  |  |  |  |  |
| 6 | Cleaning all parts for Assembly |  |  |  |  |  |  |  |
| 7 | Crystal Assembly |  |  |  |  |  |  |  |
| 8 | Preliminary Test |  |  |  |  |  |  |  |
| 9 | Spring Sealing |  |  |  |  |  |  |  |
| 10 | Pre – amplifier card assembly |  |  |  |  |  |  |  |
| 11 | Calibration with card assembly |  |  |  |  |  |  |  |
| 12 | Vibrating test at 10g for 15 mins. |  |  |  |  |  |  |  |
| 13 | Cover seating |  |  |  |  |  |  |  |
| 14 | Final Calibration |  |  |  |  |  |  |  |

Voltage Sensitivity of the Accelerometer…………………………….mV/g

Verified By:							Approved By:

Date:								(HOD, PDN)

Form No.  : R/PP-05/06/REV 00 Pg 1 of 1

ROUTINE TEST RECORD FOR LOAD / DIGITAL INDICATOR / TANK / HOPPER / SILO WEIGHING SYSTEM

SYSTEM     :

JOB No.      :						CUSTOMER:

CAPACITY:		Kg/Tonnes    	                           TAG. NO. :

| Load/Digital Indicator. | SENSOR | REMOTE (IF ANY) | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No. Serial No. | Model No. Serial No. 1) 2) 3) 4) | Model No. Serial No. | Model No. Serial No. Model No. Serial No. |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(A) -5V(A) 12V(A) -12V(A) 5V(D) 24V(I) 12V(R) |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |
| 4. | Load on the Load cell as specified in G.A. or MPL |  | Kg/T |
| 5. | Tare weight as specified in G.A. or MPL |  | Kg/T |

Resolution: L:              kg/T

Measurement to be done from 0 to 110% of the System capacity.

| Load on Load cell Kg/T or mV | Measured Load kg/T (1 to 4 Load cell system) 1 2 3 4 |  |  |  | Current output mA | Measured Load kg/T (5 to 8 Load cell system) 5 6 7 8 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Form No.  : R/PP-05/07- LI/REV 01 Pg 1 of 2

ROUTINE TEST RECORD FOR LOAD / DIGITAL INDICATOR / TANK / HOPPER / SILO WEIGHING SYSTEM

| SL. No. | Test Performed | Checked : YES/NO |
| --- | --- | --- |
| 6. | Relay Contacts: Check for the relay change over for the set points. |  |
| 7. | Display Annuciation : (As per the software Details) |  |
| 8. | RS 232 / RS 485 / BCD and Printer output (If any) The RS232 / RS485 to be checked with interfacing to PC {Attach a printout along with the test report for the Printer option } |  |
| 9. | Any other parameters (If Any) |  |
| 10. | Rating of the fuse used in the system : |  |

REMARKS (If ANY):-

Instrument Used:

Tested by:	                                                       			 Approved by:

Date:								(HOD – PDN)

Form No.  : R/PP-05/07- LI/REV 01 Pg 2 of 2

ROUTINE TEST RECORD FOR WEIGH FEEDER

JOB No.	:						CUSTOMER:

CAPACITY	:						TPH		CONVEYOR NO:

| WEIGH FEEDER | SENSOR | TACHO / PROXIMITY | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No. Serial No. LOCAL PANEL: Model No. Serial No. | Model No. Serial No. 1) 2) 3) 4) | AC / Pulse Model No. Serial No. | Model No. Serial No. Model No. Serial No. |

| Ac/Dc drive Details | Make: Capacity: Kw. Serial No. : |
| --- | --- |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(A) -5V(A) 12V(A) -12V(A) 5V(D) 24V(I) 12V(R) 15V(T) |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |
| 4. | Belt Load as specified in G.A. or MPL |  | Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL |  | m/s |

Resolution: L:              kg/m              S:                    m/s         r:                     tph           t:                        tonnes

Measurement to be done from 10 to 100% of the Feeder capacity.

I).

| Belt Length | mm. |
| --- | --- |
| Actual speed (Measured with stop clock ) | m/s |

II). Speed calibration: Calibrate the speed by entering the actual belt speed when the Drive output is 100%.

| Drive output in % | Actual Speed m/s | Measured Speed m/s |
| --- | --- | --- |
| 10 |  |  |
| 25 |  |  |
| 50 |  |  |
| 75 |  |  |
| 100 |  |  |

Form No.  : R/PP-05/07-WF/REV 00 Pg 1 of 2

ROUTINE TEST RECORD FOR WEIGH FEEDER

III). Load Calibration:   Chain Weight:		Kg/m		Cal. Rod weight:		Kg.

IV). VOLUMETRIC MODE: Control shall be less than 0.5 % of the F.S.

| Set Speed m/s |  | Measured Speed m/s |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Min | Max | Avg. | % Error |
| 10% |  |  |  |  |  |
| 25% |  |  |  |  |  |
| 50% |  |  |  |  |  |
| 75% |  |  |  |  |  |
| 100% |  |  |  |  |  |

V). GRAVIMETRIC MODE: Set rate control shall be less than 0.5% of F.S.

| Set rate tph |  | Measured rate tph Min Max Avg | Belt Load Kg/m | Totalizer ( 6 min) |  | % Error | Current O/p mA |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Act. | Meas. |  |  |
| 10% |  |  |  |  |  |  |  |
| 25% |  |  |  |  |  |  |  |
| 50% |  |  |  |  |  |  |  |
| 75% |  |  |  |  |  |  |  |
| 100% |  |  |  |  |  |  |  |

| Check for the external input response For the set rate , Digital Input and Outputs , Communication etc. |  |
| --- | --- |

Remarks:

Instrument Used:

Tested by:	  								Approved by:

Date:									(HOD – PDN)

Form No.  : R/PP-05/07-WF/REV 00 Pg 2 of 2

ROUTINE TEST RECORD FOR BELT SCALE

JOB No.	:						CUSTOMER:

CAPACITY	:						TPH		CONVEYOR NO:

| BELT SCALE | SENSOR | TACHO | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No. Serial No. REMOTE: Model No. Serial No. | Model No. Serial No. 1) 2) 3) 4) | AC / Pulse/Proximity Sensor Model No. Serial No. Rpm : N = S/3.14 * D = S Belt Speed in m/s. D Wheel dia in mts. | Model No. Serial No. Model No. Serial No. |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(A) -5V(A) 12V(A) -12V(A) 5V(D) 24V(I) 12V(R) 15V(T) |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |
| 4. | Belt Load as specified in G.A. or MPL |  | Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL |  | m/s |

Resolution: L:              kg/m              S:                    m/s         r:                     tph           t:                        tonnes

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Load on Load cell kg | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA |
| --- | --- | --- | --- | --- | --- |

Instrument Used:

Tested by:	                                                       		                         Approved by:

Date:									(HOD – PDN)

Form No.  : R/PP-05/07-BS/REV 00 Pg 1 of 2

ROUTINE TEST RECORD FOR REMOTE INDICATOR

JOB No.	:						CUSTOMER:

CAPACITY	:							 TPH:

RS 485 CONNECTION TO BE MADE AS PER WIRING DIAGRAM ENCLOSED WITH MASTER PAR LIST.

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(D) 5V (I) |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Rate in tph | Totalizer in tonnes |
| --- | --- |

Instrument Used:

Tested  by:	                                                       		                          Approved by:

Date:									 (HOD – PDN)

Form No.  : R/PP-05/07-RBS/REV 00 Pg 2 of 2

ROUTINE TEST RECORD FOR LOCAL PANEL

JOB No.	:						CUSTOMER:

CAPACITY	:						TPH

MODEL NO. :

SERIAL NO. :

RS 485 CONNECTION TO BE MADE AS PER WIRING DIAGRAM ENCLOSED WITH MASTER PAR LIST.

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(D) 5V (I) |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |
| 4. | Test for input command belt on. Belt off , Emergency off etc. |  |  |

Measurement to be done from 0 to 110% of the Full Scale capacity.

| Rate in tph | Totalizer in Tonnes |
| --- | --- |

Instrument Used:

Checked by:	                                                       		                Approved by:

Date:									   (HOD – PDN)

Form No.  : R/PP-05/07-LP/REV 00 Pg 1 of 1

ROUTINE TEST RECORD FOR WINTER

JOB No.	:						CUSTOMER:

CAPACITY	:						Kg/Tonnes	STAMP. :

| Load/Digital Indicator. | SENSOR | REMOTE (IF ANY) | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No. Serial No. | Model No. Serial No. 1) 2) 3) 4) | Model No. Serial No. | Model No. Serial No. Model No. Serial No. |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. Battery voltages (3.6 V +/- 10%) Memory card. | 5V(A) -5V(A) 12V(A) -12V(A) 5V(D) 3.6V |  |
| 3. | Check for the power on display and the Keyboard function. |  |  |
| 4. | Load on the Load cell as specified in MPL |  | Kg/T |
| 5. | Platform type (Dual or Single) |  |  |

Resolution: L:              kg/T

Measurement to be done from 0 to 100% of the System capacity.

| Load on the Load cell Kg/T or mV | Measured Load kg/T. 1 2 3 4 5 6 7 8 |
| --- | --- |

Form No.  : R/PP-05/07-WI/REV 00 Pg 1 of 2

ROUTINE TEST RECORD FOR WINTER

| SL. No. | Test Performed | Checked : YES/NO |
| --- | --- | --- |
| 6. | Relay Contacts: Check for the relay change over for the set points. |  |
| 7. | Display Annuciation : (As per the software Details) |  |
| 8. | RS 232 / RS 485 / BCD and Printer output (If ANY) {The RS232 / RS485 to be checked with interfacing to PC and use the software ProComm or software which is to be supplied} {Attach a printout along with the test report for the Printer option } |  |
| 9. | Any other parameters (If Any) |  |
| 10. | Rating of the fuse used in the system : |  |

REMARKS (If ANY):-

Instrument Used:

Checked by:	                                                        Approved by:                          (HOD – PDN)

Date:

Form No.  : R/PP-05/07-WI/REV 00 Pg 2 of 2

TEST RECORD FOR TRIP SAFE

JOB OREDR No.	:

CUSTOMER:

TRIPSAFE							LOAD CELL:

MODEL NO. :							MODEL NO. :

SERIAL NO. :							SERIAL NO  :

Calibration chart for a load input / mV input :

| SL.NO. | LOAD INPUT IN ACHIEVED |  | TRIP SAFE LEVEL ACHIEVED IN |  | FACTORY SET LEVEL | LEVEL |
| --- | --- | --- | --- | --- | --- | --- |
|  | Kg/T | % | Kg/T | % |  |  |

Percentage of error:

PF Contact for Healthy / Over Load:

Lamp indication / LED anauciation  as applicable:

Instrument Used:

Tested by:	                                                        			Approved by:

Date:									 (HOD – PDN)

Form No.  : R/PP-05/07-TS/REV 01 Pg 1 of 1

ROUTINE TEST RECORD FOR CRANE SCALE

JOB No.	:					    CUSTOMER:

CAPACITY  :                      Kg/ Tonnes	                 STAMP NO:

Load cell. Model No. :				    Serial No. :

| Crane scale | Battery charger | Hook | Bow shackle |
| --- | --- | --- | --- |
| Model No. Serial No. | ModelNo: QBC 121SR Sl.No: | Sl.No | Sl.No: |

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Battery Voltage ( For 12V ) | 10.5V to 13.6V |  |
| 2. | Derived DC Voltages (+/-5%) | 5V |  |
| 3. | Check for the power on display, Remote and the Keyboard function. |  |  |

Resolution :	L:		kg/tonnes

Measurement to be done from 0 to 100% of the Crane Scale capacity.

| Load on Load cell Kg/T or mv | Measured load Kg/T |
| --- | --- |

- Battery Charger voltage

When the battery connected:

- Low battery voltage set (CS):

- Check for the remote operation:

(For the distance of min. 10 meters)

Remote

Model No: RC500A0

Sl. No:

Calibrator

Model No: CC500A0

Sl.No:

Instrument Used:

Checked by:	 						Approved by:

Date:								(HOD – PDN)

Form No.  : R/PP-05/07-CS/REV 01 Pg 1 of 1

TEST RECORD FOR SIGNAL CONDITIONER

JOB OREDR No.	:

CUSTOMER:

SIGNAL CONDITIONER					LOAD CELL

MODEL NO. :							MODEL NO. :

SERIAL NO. :							SERIAL NO. :

Calibration chart for a load input / mV input :

| SL.NO. | LOAD ON LOAD CELL Kg/T OR mV | OUTPUT IN CURRENT / VOLTAGE |
| --- | --- | --- |

Percentage of error:

Multimeter Used:

Tested by:	                                                        			Approved by:

Date:									 (HOD – PDN)

Form No.  : R/PP-05/07-SC/REV 00 Pg 1 of 1

ROUTINE TEST RECORD FOR BATCHING SYSTEM

JOB No.	:						CUSTOMER:

CAPACITY	:						Kg/Tonnes

| BATCHING SYSTEM (BWS SERIES) | DIGITAL INDICATOR (WH) | LINE DRIVER | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No. Serial No. | Model No. Serial No. | Model No. Serial No. | Model No. Serial No. Model No. Serial No. |
| SENSOR | PLC MAKE | EXPANSION MODULE |  |
| Model No. Serial No. 1) 2) 3) 4) | Model No. Serial No. | Digital Module: QTY. : Analog Module: |  |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth :

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) The voltages are measured with respect to Respective ground. | 5V(A) -5V(A) 12V(A) -12V(A) 5V(D) 24V(I) |  |
| 3. | Check for the power on display, annunciation and the Keyboard function. |  |  |

Resolution: L:              kg/T

Measurement to be done from 0 to 110% of the Batching system capacity.

| Load on Load cell Kg/T or mV | Measured Load Kg/T 1 to 4 L/C SYS. WH1 WH2 WH3 WH4 |  |  |  | Current output mA |
| --- | --- | --- | --- | --- | --- |

Form No.  : R/PP-05/07-BCH/REV 00 Pg 1 of 2

ROUTINE TEST RECORD FOR BATCHING SYSTEM

| SL. No. | Test Performed | Checked : Yes/No |
| --- | --- | --- |
| 4. | Functional check - Manual Mode - Semi Auto Mode - Auto Mode |  |
| 5. | PF contacts to customer |  |
| 6. | PF contacts from limit switches |  |
| 7. | PF contacts for annunciation |  |
| 8. | Controls on horizontal desk |  |
| 9. | 4-20mA current output |  |
| 10. | PC software (SCADA) |  |
| 11. | Communications |  |
| 12. | Any Other Parameter(If Any) |  |
| 13. | REMARKS (IF ANY) |  |

Instrument Used:

Tested by:	  							         Approved by:

Date: 								        (HOD – PDN)

Form No.  : R/PP-05/07-BCH/REV 00 Pg 2 of 2

PERFORMANCE INDEX

| MONTH | No. OF JOBS SCHEDULED | No. OF JOBS COMPLETED | % ACHEIVED |
| --- | --- | --- | --- |
| Jun 14 Jul 14 Aug 14 Sep 14 Oct 14 Nov 14 DEC 15 JAN 15 FEB 15 MAR 15 APR 15 May 15 Jun 15 Jul 15 Aug 15 Sep 15 Oct 15 | 57 61 6567 106 80 102 72 78 73 84 66 85 85 87 76 86 | 52 65 57 94 72 93 63 65 ' 59 80 55 73 74 81 65 77 | 91.2% 91.8% 85.1% 88.7% 90% 88% 87.5% 83.3% 80.8% 95.3% 83.3% 85.9% 87.0% 93.1% 85.5% 89.5% |

Prepared  by	:

(HOD PDN)

Form No.  : R/PP-05/08/REV 01 Pg 1 of 1

ACCELEROMETER TEST RECORD

TESTED BY:								  APPROVED BY:

DATE	:						                              HOD – TRS

Form No.  : R/PP-05/09-ACC/REV 00 Pg 2 of 8

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test I Axial Sensitivity

Overall sensitivity of Std. Accelerometer and its Sig. Cond. = 43.18 mV/g

Gain of Charge amplifier (G2) = 1.022 mV/pC

Acceleration Setting Tol : +/- 10% of set value

- Frequency : 50 – 60 Hz

Acceleration: (2, 4, 8, 10, 20 & 40) * (1/9.81) g – pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

- Frequency : 50–60,100-200,300-400,700-800,1K-2K,2K-3K (Hz)

  - A= 2 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

TESTED BY:							APPROVED BY:

DATE							                 HOD – PDN

Form No.  : R/PP-05/09-ACC/REV 01 Pg 2 of 3

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test I . (b) contd.

Frequency : 50–60,100-200,300-400,700-800,1K-2K,2K-3K     (Hz)

  - A= 4 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

  - A= 8 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

  - A= 10 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

TESTED BY:							APPROVED BY:

DATE:								HOD – TRS

Form No.  : R/PP-05/09-ACC/REV 00 Pg 4 of 8

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test I . (b) contd.

Frequency : 50–60,100-200,300-400,700-800,1K-2K,2K-3K     (Hz)

  - A= 20 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

  - A= 40 * (1/9.81) g pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

Mean =           pC/g

Change sensitivity=		pC/g

TESTED BY:							APPROVED BY:

DATE							                  HOD – TRS

Form No.  : R/PP-05/09-ACC/REV 00 Pg 5 of 8

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test II Transverse Sensitivity

Overall sensitivity of Std, Accelerometer and its Sig. Cond. = 50.102 mV/g

Gain of Charge amplifier (G2) = 1.022 mV/pC

Acceleration Setting Tol : +/- 10% of set value

- A= 2 * (1/9.81) g pk-pk 			Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

- A= 4 * (1/9.81) g pk-pk 			Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

- A= 8 * (1/9.81) g pk-pk			Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

TESTED BY:							APPROVED BY:

DATE								  HOD – TRS

Form No.  : R/PP-05/09-ACC/REV 00 Pg 6 of 8

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test II Contd.

- A= 10 * (1/9.81) g pk-pk 		Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

- A= 20 * (1/9.81) g pk-pk 		Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

- A= 40 * (1/9.81) g pk-pk			Frequency shall be 50 +/- 2 Hz

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/50.102 | Sensitivity S= (V21)/A pC/g | %= (SX100)/ Avg. ‘S’ |
| --- | --- | --- | --- | --- | --- | --- |

Mean =         %

Transverse sensitivity=		%

TESTED BY:							APPROVED BY:

DATE								 HOD – TRS

Form No.  : R/PP-05/09-ACC/REV 00 Pg 7 of 8

ACCELEROMETER TEST RECORD

Model No……………… Sl.No………………………..

Test III Post Sealing Axial Sensitivity

Overall sens. of Std. Accelerometer and its Sig. Cond.(G1) =  43.18   mV/g

Gain of Charge amplifier (G2) = 1.022 mV

Acceleration Setting Tol : +/- 10% of set value

Frequency : 50 – 60 Hz

Acceleration: (2, 4, 8, 10, 20 & 40) * (1/9.81) g – pk-pk

| Freq. (Hz) | O/P of std. Accl. V1(pk-pk)mV | O/P of test Accl. V2(pk-pk)mV | V21 = (V2/G2) mV | Acceleration A= (V1)/G1 | Sensitivity S= (V21)/A pC/g |
| --- | --- | --- | --- | --- | --- |

TESTED BY:							APPROVED BY:

DATE:								 HOD – (PDN)

Form No.  : R/PP-05/09-ACC/REV 01 Pg 3 of 3

VIBRATION METER TEST RECORD

Model No……………… Sl.No………………………..

| POINTS | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| EXPECTED DC VOLTAGE (V) | 1<=V<=4 | 1<=V<=4 | ± 1 V | ± 20 mV | ± 2 mV | ± 5 mV | ± 10 mV | ± 10 mV | ± 1 mV | ± 2 mV |
| OBTAINED DC VOLTAGE (V) |  |  |  |  |  |  |  |  |  |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/10-VM/REV 00 Pg 2 of 6

VIBRATION METER TEST RECORD

Model No……………… Sl.No………………………..

| POINTS | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| EXPECTED AC VOLTAGE ± 5% | 2 V p-p | 10V p-p | 3V p-p | 0.5V p-p | 1.5 V p-p | 0.6 V 1.8V | 1 V DC | 3 V DC |
| OBTAINED AC VOLTAGE |  |  |  |  |  |  |  |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/10-VM/REV 00 Pg 3 of 6

VIBRATION METER TEST RECORD

Model No……………… Sl.No………………………..

| Freq. Hz | Accl. (peak m/sec2 ) | Vel.(rms) Cm/sec Cal.val. | Vel.(rms) Cm/sec Meas.val. | % Error | Disp mm pk-pk cal.val. | Disp mm pk-pk Meas.val. | % Error |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |
| 50 |  |  |  |  |  |  |  |
| 71.17 |  |  |  |  |  |  |  |
| 100 |  |  |  |  |  |  |  |
| 112.5 |  |  |  |  |  |  |  |
| 200 |  |  |  |  |  |  |  |
| 500 |  |  |  |  |  |  |  |
| 1K |  |  |  |  |  |  |  |
| 2K |  |  |  |  |  |  |  |
| 3K |  |  |  |  |  |  |  |
| 5K |  |  |  |  |  |  |  |
| 10K |  |  |  |  |  |  |  |
| 20K |  |  |  |  |  |  |  |
| 100K |  |  |  |  |  |  |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/10-VM/REV 00 Pg 4 of 6

VIBRATION METER TEST RECORD

Model No……………… Sl.No………………………..

| Freq. Hz | Accl. (peak m/sec2 ) | Vel.(rms) Cm/sec Cal.val. | Vel.(rms) Cm/sec Meas.val. | % Error | Disp mm pk-pk cal.val. | Disp mm pk-pk Meas.val. | % Error |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 50 |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |

| Freq. Hz | Accl. (peak m/sec2 ) | Disp mm pk-pk cal.val. | Disp mm pk-pk Meas.val. | % Error |
| --- | --- | --- | --- | --- |
| 50 |  |  |  |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Name				Signature			Designation

Form No.  : R/PP-05/10-VM/REV 00 Pg 5 of 6

VIBRATION METER TEST RECORD

Model No……………… Sl.No………………………..

| VELOCITY |  |  |  |  | ACCELERATION |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Freq. Hz | O/p (V1) p-pmV | Cal.val. (rms) Cm/sec | Meas.val. (rms) Cm/sec | % Error | Freq. Hz | O/p (V1) p-pmV | Cal.val. (pk) m/sec2 | Meas.val. (pk) m/sec2 | % Error |

| DISPLACEMENT |  |  |  |  | DISPLACEMENT |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Freq. Hz | O/p (V1) p-pmV | Cal.val. (p-p) mm | Meas.val. (p-p) mm | % Error | Freq. Hz | O/p (V1) p-pmV | Cal.val. (p-p) mm | Meas.val. (p-p) mm | % Error |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/10-VM/REV 00 Pg 6 of 6

CHARGE AMPLIFIER TEST RECORD

Customer……………………………………………………………………………… J.O. No……………………………

Model No……………… Sl.No…………………………

| POINTS | TP4 | TP5 | TP6 | TP1 | TP2 | TP7 |
| --- | --- | --- | --- | --- | --- | --- |
| EXPECTED DC VOLTAGE (V) | 6.2 to 7.5 | 1<=V<=4 | 1<=V<=4 | +6V to +6.5 | -6V to -6.5 | Max+1V Min -1V |
| OBTAINED DC VOLTAGE (V) |  |  |  |  |  |  |

|  | EXPECTED CURRENT | OBTAINED CURRENT |
| --- | --- | --- |
| With i/p open with 10K load | 3mA± 1mA |  |
| With i/p 10V (p-p) .1 KHz AC Signal With 10K load | 3mA± 1mA |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/11-CA/REV 00 Pg 2 of 7

CHARGE AMPLIFIER TEST RECORD

Customer……………………………………………………………………………… J.O. No……………………………

Model No……………… Sl.No…………………………

READING TOL. +/- 2%

a)	10 Hz						b)	20Hz

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/11-CA/REV 00 Pg 3 of 7

CHARGE AMPLIFIER TEST RECORD

Model No……………… Sl.No…………………………

Input Signal Frequency:

c)	50 Hz						d)	100Hz

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

e)	200 Hz						f)	500Hz

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/11-CA/REV 00 Pg 4 of 7

CHARGE AMPLIFIER TEST RECORD

Model No……………… Sl.No…………………………

Input Signal Frequency:

g)	1K Hz						h)	2K Hz

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

i)	5K Hz						j)	10K Hz

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

| SL. No. | I/P Signal pk-pk | O/P Signal Pk-pk |
| --- | --- | --- |
| 1 | 10mV |  |
| 2 | 20mV |  |
| 3 | 50mV |  |
| 4 | 100mV |  |
| 5 | 200mV |  |
| 6 | 500mV |  |
| 7 | 1V |  |
| 8 | 2V |  |
| 9 | 5V |  |
| 10 | 10V |  |

TESTED BY:					APPROVED BY:                             HOD – PDN

DATE

Form No.  : R/PP-05/11-CA/REV 00 Pg 5 of 7

CHARGE AMPLIFIER TEST RECORD

Model No……………… Sl.No…………………………

TEST IV

Measurement of Maximum Input Signal and

Maximum Output Signal

Maximum Input Signal = ………………………………………………Freq……………………

Maximum Input Signal = ………………………………………………Freq……………………

TEST V Frequency Response of Charge Amplifier

(A)	Input Signal = 50 mV pk-pk

READING TOL. ± 2%

| Freq. Hz. | Input Signal Vi mV pk-pk | Output Signal Vo mV pk-pk | Gain = 20 log (Vo/Vi) |
| --- | --- | --- | --- |
| 1 | 50 |  |  |
| 2 | 50 |  |  |
| 5 | 50 |  |  |
| 10 | 50 |  |  |
| 20 | 50 |  |  |
| 50 | 50 |  |  |
| 100 | 50 |  |  |
| 200 | 50 |  |  |
| 500 | 50 |  |  |
| 1K | 50 |  |  |
| 2K | 50 |  |  |
| 5K | 50 |  |  |
| 10K | 50 |  |  |
| 20K | 50 |  |  |
| 50K | 50 |  |  |
| 100K | 50 |  |  |

TESTED BY:					APPROVED BY:

DATE						 HOD – PDN

Form No.  : R/PP-05/11-CA/REV 00 Pg 6 of 7

CHARGE AMPLIFIER TEST RECORD

Model No……………… Sl.No…………………………

TEST V (CONTD.)

(B)	Input Signal = 10 V pk-pk

READING TOL. ± 2%

| Freq. Hz. | Input Signal Vi mV pk-pk | Output Signal Vo mV pk-pk | Gain = 20 log (Vo/Vi) |
| --- | --- | --- | --- |
| 1 | 10 |  |  |
| 2 | 10 |  |  |
| 5 | 10 |  |  |
| 10 | 10 |  |  |
| 20 | 10 |  |  |
| 50 | 10 |  |  |
| 100 | 10 |  |  |
| 200 | 10 |  |  |
| 500 | 10 |  |  |
| 1K | 10 |  |  |
| 2K | 10 |  |  |
| 5K | 10 |  |  |
| 10K | 10 |  |  |
| 20K | 10 |  |  |
| 50K | 10 |  |  |
| 100K | 10 |  |  |

TEST VI Measurement of Noise Level

Accelerometer type………………………… Sl No……………………………..

| NOISE LEVEL |  |  |
| --- | --- | --- |

TESTED BY:					APPROVED BY:                    HOD – PDN

DATE:

Customer			 				J.O. No.

Form No.  : R/PP-05/11-CA/REV 00 Pg 7 of 7

CHARGE AMPLIFIER TEST RECORD

Customer……………………………………………………………………………… J.O. No……………………………

Model No……………… Sl.No…………………………

SPECIFICATION:

A)	FREQUENCY RANGE		:

B)	MAXIMUM INPUT SIGNAL	:

C)	MAXIMUM OUTPUT SIGNAL	:

D)	GAIN				:

- DIMENSIONS

HEIGHT			:

WIDTH				:

DEPTH				:

TESTED BY:								APPROVED BY:

DATE									HOD – PDN

Name				Signature			Designation

Form No.  : R/PP-05/11-CA/REV 00 Pg 1 of 7

VIBRATION METER TEST RECORD

Customer……………………………………………………………………………… J.O. No……………………………

Model No……………… Sl.No…………………………

ACCELEROMETER MODEL NO……………………………………… SL No…………………………………………………

ACCURACY

ACCELERATION			:

VELOCITY				:

DISPLACEMENT			:

TESTED BY:								APPROVED BY:

DATE									HOD – PDN

Form No.  : R/PP-05/10-VM/REV 00 Pg 1 of 6

ACCELEROMETER TEST RECORD

Customer……………………………………………………………………………… J.O. No……………………………

Model No……………… Sl.No…………………………

SPECIFICATION:

A)	CHARGE SENSITIVITY		:

B)	CAPACITANCE			:

C)	LEAKGE RESISTANCE		:

At 30oC , M ohms

D)	TRANSVERSE SENSITIVITY	:

- CONSTRUCTION			:	CENTER MOUNTED

COMPRESSION

- TYPE OF CONNECTION		: 	SIDE

- OUTPUT CONNECTOR			:	10-32	UNF

- WEIGHT (gms)			:

- MOUNTING THREAD			:	M5

- HEIGHT				:

- SPANNER SIZE			:

ACCESSORIES STANDARD

LOW NOISE CABLE OF 1 METER LENGTH WITH

10-32 UNF MINIATURE COAXIAL MALE CONNECTOR (	)………………1 No……………

THREADED STEEL STUDS M5 (AS 1001)…………………………………5 NOS.

TESTED BY:					                            APPROVED BY:

DATE								         HOD – PDN

Form No.  : R/PP-05/09-ACC/REV 01 Pg 1 of 3

PERFORMANCE DELAY ANALYSIS

(Production Systems/Production TRS)

| MONTH | No. OF JOBS / LOAD CELLS SHEDULED | No. OF JOBS / LOAD CELLS COMPLETED | PERCENTAGE OF DELAY DUE TO |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nov 15 | 62 | 54 | REWORK & REJECTIONS IN | MAN POWER | COMPONENT / MATERIAL STORAGE | PROLONG TESTING | DESIGN I/P |  |
|  |  |  | Nil | Nil | 8.0% | 3.2% | 1.6% |  |
| PREPARED BY: APPROVED BY: HOD : PDN |  |  |  |  |  |  |  |  |
| Form No. : R/PP-05/12/REV 01 Pg 1 of 2 |  |  |  |  |  |  |  |  |

PARTICULARS FOR DELAY ANALYSIS (Production Systems/Production TRS)

MONTH / YEAR:  OCT 15

| SL.NO. | J.O.NO. | CUSTOMER | DEPARTMENT | REASON |
| --- | --- | --- | --- | --- |
| 01 02 03 04 05 06 07 08 09 10 11 12 | 405/15 561/15 662/15 359/15 721/15 710/15 590/15 626/15 701/15 710/15 126/15 531/15 | Chemical Constrn Coromandel Int. Electromech Matl. Graduate Agro Pragati Automation Ducom Thyssenkrupp India UCIL MPRL Thyssenkrupp AG Conveying Steed Engg. | PDN PDN Sub cont PUR DSN PDN PDN PDN PUR Sub Cont PUR DSN | Delay in Mechanical Delay in Mechanical Delay in Load cell blank Delay in procurement of communication card Delay in Design Delay in testing Delay in Load cell Delay in testing Delay in Load cell Delay in enclosure Delay in Belt procurement Delay design |

TOTAL NO OF JOBS SCHEDULED :     95 Jobs

NO. OF JOBS EXECUTED	        :     83 Jobs

TOTAL DELAY                                     :    12.6%

PREPARED BY: Arul				  APPROVED BY:

DATE:	05/ 10/ 15						            	                    	(HOD  PDN)

COPY:	1)   CEO

2    MR

3)   Planning

Form No.  : R/PP-05/12/REV 01 Pg 2 of 2

REWORK DUE TO DESIGN INPUT (Production Systems)

MONTH/YEAR:

| SL.No. | J.O.No. | CUSTOMER | SYSTEM / MATERIAL | REASON FOR REWORK | QTY | TIME TAKEN |
| --- | --- | --- | --- | --- | --- | --- |

TOTAL No. OF JOBS SCHEDULED	:

No. OF JOBS EXECUTED		:

TOTAL DELAY	 in %			:

PREPARED BY:							APPROVED BY:

DATE:									(  H.O.D PDN)

COPY:	1)	MR

- DSN

Form No.  : R/PP-05/13/REV 01 Pg 1 of 1

INCOMING MATERIAL REVIEW OF PDN

MONTH: Dec 2012

| SL.NO. | JOB No. | CUSTOMER | MAT. DESCRIPTION | NATURE OF NON- CONFORMANCE | REMARKS |
| --- | --- | --- | --- | --- | --- |
| 01. |  |  |  |  |  |

COPY TO: 1) Director Technical

2) MR

3) QC

PREPARED BY:								APPROVED BY

HOD(PDN)

Form No.  : R/PP-05/14/REV 00 Pg 1 of 1

TEST REPORT FOR VIBRATION METER

JOB NO :                                                                                             ACCELEROMETER

CUSTOMER:                                                                                       MODEL NO:

VIBRATION METER                                                                         SERIAL NO:

MODEL NO:

SERIAL NO:

| ACCELERATION in m/s^2 | VELOCITY in cm/s | DISPLACEMENT in mm |
| --- | --- | --- |
| STANDARD TESTED | STANDARD TESTED | STANDARD TESTED |

TESTED BY:                                                                           APPROVED BY:

DATE:                                                                                          (HOD-PDN)

Note:  This format shall be used when testing and calibration is done using vibration meter standard reference.

Form No: R/PP-05/10-VM –I Rev -00

Test Report for miscellaneous items: Production Systems

Job No:                                                                              Model No:

Customer:                                                                         Sl. No:

Item  Description :                                                           Qty:

_______________________________________________________________________________________

Test Readings and Result:

Tested by:                                                                         	  Approved by:

Date:                                                                                          (HOD-PDN)

Form No.: R/PP-05/16 Rev 00 Page 1 of  1

PRODUCTION (Transducer)

DETAIL PRODUCTION PLAN: MONTH:

| J.O.No. | Customer | Model | QTY. | DATE |  | REMARKS |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | PLANNED | ACTUAL |  |

PREPARED BY:												APPROVED BY:

(PDN-IN-CHARGE)												(HOD - PDN)

Form No.  : R/PP-05/02/REV 00 Pg 1 of 1

| TRANSDUCER PRODUCTION PERFORMANCE INDEX YEAR- |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| MONTH | No. OF LOAD CELLS SCHEDULED | LOAD CELL PRODUCED |  |  | LOAD CELLS REJECTED |  |
|  |  | TOTAL QTY | EXCESS % | SHORTAGE % | TOTAL QTY | % REJECTION |
| JAN |  |  |  |  |  |  |
| FEB |  |  |  |  |  |  |
| MAR |  |  |  |  |  |  |
| APR |  |  |  |  |  |  |
| MAY |  |  |  |  |  |  |
| JUN |  |  |  |  |  |  |
| JUL |  |  |  |  |  |  |
| AUG |  |  |  |  |  |  |
| SEP |  |  |  |  |  |  |
| OCT |  |  |  |  |  |  |
| NOV |  |  |  |  |  |  |
| DEC |  |  |  |  |  |  |
| PREPARED BY: APPROVED BY: IN CHARGE TRS. (HOD - PDN) |  |  |  |  |  |  |

Form No.  : R/PP-05/15/REV 00 Pg 1 of 1

TEST REPORT FOR ACCELEROMETER WITH  4-20 mA OUTPUT

CUSTOMER:                                                                                                                       MODEL No:

JOB No.:                                                                                                                               SL No:

1. LIBEARITY TEST: (Ref Freq: 100 Hz)

| SL.No | ACCELERATION in “ g” | CURRENT OUTPUT in mA |
| --- | --- | --- |

2. FREQUENCY RESPONSE TEST :

| SL.No | FREQUENCY in Hz | ACCELERATION in “g” | CURRENT OUTPUT in mA |
| --- | --- | --- | --- |

TESTED BY :                                                                                                                                         APPROVED BY:

DATE          :                                                                                                                                          (H.O.D, PDN)