ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : BS-COMPLETE-VERIFY					CUSTOMER NAME: ABEL STEEL CORP

CAPACITY : 100					CONVEYOR NO: BC-01

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-100 Serial No: BS-SN-001 S/W Version: 3.1.3 REMOTE: Model No: RM-100 Serial No: RM-SN-001 | Model No: LC-50 Serial No: 1) LC-01 2) LC-02 3) LC-03 4) LC-04 | Model No: TS-10 Serial No: TS-SN-01 | Model No: JB-1 Serial No: JB-SN-1 Model No: JB-2 Serial No: JB-SN-2 Model No: JB-3 Serial No: JB-SN-3 |
|  |  | ANGLE SENSOR |  |
|  |  | Model No: AS-300 Serial No: AS-SN-101 |  |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth : OK (100M Ohm)

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | Display OK | Display OK |
| 4. | Belt Load as specified in G.A. or MPL | 27.78 | 27.78 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 2.50 | 2.50 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | Wiring OK | Wiring OK |

Resolution: L: 0.01 kg/m          S: 0.01 m/s          r: 0.1 tph          t: 0.1 tonnes

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Load on Load cell Kg/Mv | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA |
| --- | --- | --- | --- | --- | --- |
| 10.0 | 2.50 | 2.50 | 10.0 | 1.5 | 2.00 |
| 20.0 | 5.00 | 2.50 | 20.0 | 3.0 | 4.00 |
| 30.0 | 7.50 | 2.50 | 30.0 | 4.5 | 6.00 |
| 40.0 | 10.00 | 2.50 | 40.0 | 6.0 | 8.00 |
| 50.0 | 12.50 | 2.50 | 50.0 | 7.5 | 10.00 |
| 60.0 | 15.00 | 2.50 | 60.0 | 9.0 | 12.00 |
| 70.0 | 17.50 | 2.50 | 70.0 | 10.5 | 14.00 |
| 80.0 | 20.00 | 2.50 | 80.0 | 12.0 | 16.00 |
| 90.0 | 22.50 | 2.50 | 90.0 | 13.5 | 18.00 |
| 100.0 | 25.00 | 2.50 | 100.0 | 15.0 | 20.00 |
| 110.0 | 27.50 | 2.50 | 110.0 | 16.5 | 22.00 |

Instrument Used: Fluke Multimeter

Tested by: System Admin					Approved by: HOD-PDN

Date: 14-Sep-2026						(HOD - PDN)