ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : BS-FINAL-1PAGE-TEST						CUSTOMER NAME: ABEL STEEL CORP

CAPACITY : 1500									CONVEYOR NO: CV-01

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-2000 Serial No: BS-SN-001 REMOTE: Model No: RM-2000 Serial No: RM-SN-001 | Model No: BR3200H0 Serial No: 1) SN1 2) SN2 3) SN3 4) SN4 | Sensor: Model No: TS-100 Serial No: TS-SN1 | Model No: JB-1 Serial No: JB-SN1 Model No: Serial No: |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth :

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | OK | OK |
| 4. | Belt Load as specified in G.A. or MPL | 208.33 | 208.33 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 2.00 | 2.00 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | OK | OK |

Resolution                       L: 0.1                          S: 0.01                                 r: 1                           t: 0.1

| Load on Load cell Kg / mV | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA AO1 | O/p current mA AO2 |
| --- | --- | --- | --- | --- | --- | --- |
| 3 | 0.0 | 2.00 | 0 | 0.0 | 4.00 | 4.00 |
| 3 | 41.7 | 2.00 | 300 | 30.0 | 7.20 | 7.20 |
| 4 | 83.3 | 2.00 | 600 | 60.0 | 10.40 | 10.40 |
| 4 | 125.0 | 2.00 | 900 | 90.0 | 13.60 | 13.60 |
| 5 | 166.7 | 2.00 | 1200 | 120.0 | 16.80 | 16.80 |
| 5 | 208.3 | 2.00 | 1500 | 150.0 | 20.00 | 20.00 |

Instrument Used: 	Fluke 87V Multimeter & 50kg Class F Weights

Tested by: System Admin	                                                       		                         Approved by:

Date: 14-Sep-2026									            (HOD - PDN)