ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : BS-ORIGINAL-TEMPLATE-FINAL						CUSTOMER NAME: ABEL STEEL CORP

CAPACITY : 1000 TPH									CONVEYOR NO: BC-01

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-10000 Serial No: BS-SN-001 REMOTE: Model No: RM-100 Serial No: RM-SN-001 | Model No: LC-500 Serial No: 1) LC-01 2) LC-02 3) LC-03 4) LC-04 | Sensor: Model No: SS-100 Serial No: TS-SN-01 | Model No: JB-101 Serial No: JB-SN-101 Model No: JB-102 Serial No: JB-SN-102 |

Resolution:    L: 0.1 kg/m          S: 0.01 m/s         r: 1 tph           t: 0.1 tonnes

Check for short between Line & Neutral, Neutral & Earth, Line & Earth : OK (100M Ohm)

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | Display OK | Display OK |
| 4. | Belt Load as specified in G.A. or MPL | 27.78 | 27.78 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 2.50 | 2.50 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | Wiring OK | Wiring OK |

Resolution                       L: 0.1                          S: 0.01                                 r: 1                           t: 0.1

| Load on Load cell Kg / mV | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA AO1 | O/p current mA AO2 |
| --- | --- | --- | --- | --- | --- | --- |
| 10 | 2.5 | 2.50 | 10 | 1.5 | 2.00 | 2.00 |
| 20 | 5.0 | 2.50 | 20 | 3.0 | 4.00 | 4.00 |
| 30 | 7.5 | 2.50 | 30 | 4.5 | 6.00 | 6.00 |
| 40 | 10.0 | 2.50 | 40 | 6.0 | 8.00 | 8.00 |
| 50 | 12.5 | 2.50 | 50 | 7.5 | 10.00 | 10.00 |
| 60 | 15.0 | 2.50 | 60 | 9.0 | 12.00 | 12.00 |
| 70 | 17.5 | 2.50 | 70 | 10.5 | 14.00 | 14.00 |
| 80 | 20.0 | 2.50 | 80 | 12.0 | 16.00 | 16.00 |
| 90 | 22.5 | 2.50 | 90 | 13.5 | 18.00 | 18.00 |
| 100 | 25.0 | 2.50 | 100 | 15.0 | 20.00 | 20.00 |
| 110 | 27.5 | 2.50 | 110 | 16.5 | 22.00 | 22.00 |

Instrument Used: 	Fluke 87V Multimeter

Tested by: System Admin	                                                       		                         Approved by: HOD-PDN

Date: 14-Sep-2026									            (HOD - PDN)