ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : BS-10001	CUSTOMER NAME: ABEL STEEL CORP

CAPACITY : 1000 TPH	CONVEYOR NO: BC-01

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-10000 Serial No: BS-SN-001 S/W Version: 3.1.3 REMOTE: Model No: RM-100 Serial No: RM-SN-001 | Model No: LC-500 Serial No: 1) LC-01 2) LC-02 3) LC-03 4) LC-04 | Model No: SS-100 Serial No: SS-SN-01 | Model No. : JB-101 Serial No. : JB-SN-101 Model No. : JB-102 Serial No. : JB-SN-102 Model No. : JB-103 Serial No. : JB-SN-103 |
|  |  | ANGLE SENSOR |  |
|  |  | Model No: AS-300 Serial No: AS-SN-101 |  |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth : OK (100M Ohm)

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) |  |  |
| 2. | Derived DC Voltages (+/-5%) |  |  |
| 3. | Display and Keypad Functionality |  |  |
| 4. | Belt Load as specified in G.A. or MPL | 277.8 | 277.8 277.8 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 1.00 | 1.00 1.00 m/s |
| 6. | Number of PF Contacts |  |  |
| 7. | Communication Output (Specify) |  |  |
| 8. | Analog Output (Specify number) |  |  |
| 9. | Wiring, TB and Component Layout |  |  |

Resolution:    L: 0.1 kg/m          S: 0.01 m/s         r: 1 tph           t: 0.1 tonnes

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Load on Load cell Kg/Mv | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA |
| --- | --- | --- | --- | --- | --- |

Instrument Used: Fluke 87V Multimeter

Tested by: Senior Test Engineer								Approved by: HOD-PDN

Date: 14-Sep-2026									(HOD - PDN)