ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : BS-TEST-UI-999						CUSTOMER NAME:

CAPACITY : 1200 TPH									CONVEYOR NO: BC-05

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-1200 Serial No: BS-SN-999 S/W Version: REMOTE: Model No: Serial No: | Model No: Serial No: 1) LC-750-1 2) LC-750-2 3) LC-750-3 4) LC-750-4 | Model No: Serial No: | Model No. : Serial No. : Model No. : Serial No. : Model No. : Serial No. : |
|  |  | ANGLE SENSOR |  |
|  |  | Model No: Serial No: |  |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth :

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | Display OK | Display OK |
| 4. | Belt Load as specified in G.A. or MPL | 27.78 | 27.78 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 2.50 | 2.50 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | Wiring OK | Wiring OK |

Resolution:    L: 0.1 kg/m          S: 0.01 m/s         r: 1 tph           t: 0.1 tonnes

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Load on Load cell Kg/Mv | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA |
| --- | --- | --- | --- | --- | --- |
| 10 | 2.5 | 2.50 | 10 | 1.5 | 2.00 |
| 20 | 5.0 | 2.50 | 20 | 3.0 | 4.00 |
| 30 | 7.5 | 2.50 | 30 | 4.5 | 6.00 |
| 40 | 10.0 | 2.50 | 40 | 6.0 | 8.00 |
| 50 | 12.5 | 2.50 | 50 | 7.5 | 10.00 |
| 60 | 15.0 | 2.50 | 60 | 9.0 | 12.00 |
| 70 | 17.5 | 2.50 | 70 | 10.5 | 14.00 |
| 80 | 20.0 | 2.50 | 80 | 12.0 | 16.00 |
| 90 | 22.5 | 2.50 | 90 | 13.5 | 18.00 |
| 100 | 25.0 | 2.50 | 100 | 15.0 | 20.00 |
| 110 | 27.5 | 2.50 | 110 | 16.5 | 22.00 |

Instrument Used: Fluke Multimeter

Tested by: System Admin								Approved by: HOD-PDN

Date: 14-Sep-2026									(HOD - PDN)