ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : 						CUSTOMER NAME:

CAPACITY : 									CONVEYOR NO:

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: Serial No: REMOTE: Model No: Serial No: | Model No: Serial No: 1) 2) 3) 4) | Sensor: Proximity Model No: Serial No: | Model No: Serial No: Model No: Serial No: |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth :

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | OK | OK |
| 4. | Belt Load as specified in G.A. or MPL | 5.00 | 27.78 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 23.87 | 1.00 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | OK | OK |

Resolution                       L: 0.01                          S: 0.01                                 r: 0.1                           t: 0.1

| Load on Load cell Kg / mV | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA AO1 | O/p current mA AO2 |
| --- | --- | --- | --- | --- | --- | --- |
| 3.0 | 0.00 | 1.00 | 0.0 | 0.0 | 4.00 | 4.00 |
| 3.4 | 5.56 | 1.00 | 20.0 | 2.0 | 7.20 | 7.20 |
| 3.8 | 11.11 | 1.00 | 40.0 | 4.0 | 10.40 | 10.40 |
| 4.2 | 16.67 | 1.00 | 60.0 | 6.0 | 13.60 | 13.60 |
| 4.6 | 22.22 | 1.00 | 80.0 | 8.0 | 16.80 | 16.80 |
| 5.0 | 27.78 | 1.00 | 100.0 | 10.0 | 20.00 | 20.00 |

Instrument Used:

Tested by: System Admin	                                                       		                         Approved by:

Date: 14-Sep-2026									            (HOD - PDN)