ROUTINE TEST RECORD FOR BELT SCALE

JOB No. : JOB-2026-812						CUSTOMER NAME: Tata Steel Mining

CAPACITY : 1000									CONVEYOR NO: CV-39B

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BS-MODEL-812 Serial No: SN-BS-7812 REMOTE: Model No: RM-UNIT-01 Serial No: SN-RM-5812 | Model No: LC-250KG Serial No: 1) LC1-002 2) LC2-002 3) LC3-002 4) LC4-002 | Sensor: Proximity Model No: TACHO-SENS-PRO Serial No: TS-8812 | Model No: JBOX-IP67 Serial No: JB-SN-912 Model No: JBOX-IP67 Serial No: JB-SN-1012 |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth : OK (Checked Line/Neutral, Neutral/Earth, Line/Earth)

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Display and Keypad Functionality | OK | OK |
| 4. | Belt Load as specified in G.A. or MPL | 141.0 kg/m | 141.0 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 1.97 | 1.97 m/s |
| 6. | Number of PF Contacts | 4 | 4 |
| 7. | Communication Output (Specify) | RS-485 | RS-485 |
| 8. | Analog Output (Specify number) | 4-20mA | 4-20mA |
| 9. | Wiring, TB and Component Layout | OK | OK |

Resolution                       L: 0.1                          S: 0.01                                 r: 1                           t: 0.1

| Load on Load cell Kg / mV | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA AO1 | O/p current mA AO2 |
| --- | --- | --- | --- | --- | --- | --- |
| 3 | 0.0 | 1.97 | 0 | 0.0 | 4.00 | 4.00 |
| 3 | 28.2 | 1.97 | 200 | 20.0 | 7.20 | 7.20 |
| 4 | 56.4 | 1.97 | 400 | 40.0 | 10.40 | 10.40 |
| 4 | 84.6 | 1.97 | 600 | 60.0 | 13.60 | 13.60 |
| 5 | 112.8 | 1.97 | 800 | 80.0 | 16.80 | 16.80 |
| 5 | 141.0 | 1.97 | 1000 | 100.0 | 20.00 | 20.00 |

Instrument Used: 	Fluke 87V Digital Multimeter & 50kg Calibrated Class F Weights

Tested by: A	                                                       		                         Approved by: HOD-PDN

Date: 14-Sep-2026									            (HOD - PDN)