TEST RECORD FOR SIGNAL CONDITIONER

JOB ORDER No:	JOB-2026-691

CUSTOMER:	Adani Power Thermal

SIGNAL CONDITIONER	LOAD CELL

MODEL NO:	SC-DIN-420	MODEL NO:	LC-SHEAR-3000

SERIAL NO:	SC-2026-001	SERIAL NO:	LC-88901

Linearity Test

| SL.NO. | LOAD ON LOAD CELL Kg/T OR mV | OUTPUT IN CURRENT / VOLTAGE |
| --- | --- | --- |
| 1 | 0.0 mV | 4.00 mA |
| 2 | 5.0 mV | 7.20 mA |
| 3 | 10.0 mV | 10.40 mA |
| 4 | 15.0 mV | 13.60 mA |
| 5 | 20.0 mV | 16.80 mA |
| 6 | 25.0 mV | 20.00 mA |

Check the POWER ON Annunciation:	Checked & Found OK

Percentage of error:	0.02 %

Multimeter Used:	Fluke 87V (Sl No: FLK-9901)

Tested by:	A	Approved by:	HOD-PDN

Date:	17-Sep-2026	(HOD - PDN)