ROUTINE TEST RECORD FOR LOAD INDICATOR / DIGITAL INDICATOR / TANK / HOPPER / SILO WEIGHING SYSTEM/CRANE WEIGHING SYSTEM

SYSTEM  : DIGITAL INDICATOR

JOB No.    : JOB-2026-857			CUSTOMER NAME: Adani Power Thermal

CAPACITY: 1000 Kg	                                       TAG. NO: CV-71B

| Load/Digital Indicator. | SENSOR | REMOTE (IF ANY) | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: DI-3000 Serial No: SN-DI-8821 S/W Version: v3.2.0 | Model No: LC-CZ50T Serial No: 1) LC1-007 5) LC-SN-005 2) LC2-007 6) LC-SN-006 3) LC3-007 7) LC-SN-007 4) LC4-007 8) LC-SN-008 | Model No: N/A Serial No: N/A | Model No: JB-8P Serial No: JB-SN-441 |

Check for short between Line & Neutral , Neutral & Earth , Line & Earth : OK (Checked Line/Neutral, Neutral/Earth, Line/Earth)

| SL. No. | Test Performed | Specified Values | Actual Values |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V | 230V |
| 2. | Derived DC Voltages (+/-5%) | 24V | 24V |
| 3. | Check for the power on display and the Keyboard function. | OK | OK |
| 4. | Load on the Load cell as specified in G.A. or MPL | Full Capacity Span | Passed |
| 5. | Tare weight as specified in G.A. or MPL | Zero Balance OK | Passed |

Resolution: L: kg/T

Measurement to be done from 0 to 110% of the System capacity.

| Load on Load cell Kg/T or mV | Measured Load kg/T (1 to 4 Load cell system) 1 2 3 4 |  |  |  | Current output mA | Measured Load kg/T (5 to 8 Load cell system) 5 6 7 8 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3.00 | 0.0 | 2.63 | 0 | 0.0 | 4.00 | 0 | 0 | 0 | 0 |
| 3.40 | 21.1 | 2.63 | 200 | 20.0 | 7.20 | 100 | 100 | 100 | 100 |
| 3.80 | 42.2 | 2.63 | 400 | 40.0 | 10.40 | 200 | 200 | 200 | 200 |
| 4.20 | 63.4 | 2.63 | 600 | 60.0 | 13.60 | 300 | 300 | 300 | 300 |
| 4.60 | 84.5 | 2.63 | 800 | 80.0 | 16.80 | 400 | 400 | 400 | 400 |
| 5.00 | 105.6 | 2.63 | 1000 | 100.0 | 20.00 | 500 | 500 | 500 | 500 |
| 60% | 600.0 | 600.0 | 4.20 | 13.60 | 13.60 | 600 | 600 | 600 | 600 |
| 70% | 700.0 | 700.0 | 4.40 | 15.20 | 15.20 | 700 | 700 | 700 | 700 |
| 80% | 800.0 | 800.0 | 4.60 | 16.80 | 16.80 | 800 | 800 | 800 | 800 |
| 90% | 900.0 | 900.0 | 4.80 | 18.40 | 18.40 | 900 | 900 | 900 | 900 |
| 100% | 1000.0 | 1000.0 | 5.00 | 20.00 | 20.00 | 1000 | 1000 | 1000 | 1000 |
| 110% | 1100.0 | 1100.0 | 5.20 | 21.60 | 21.60 | 1100 | 1100 | 1100 | 1100 |

ROUTINE TEST RECORD FOR LOAD / DIGITAL INDICATOR / TANK / HOPPER / SILO WEIGHING SYSTEM

| SL. No. | Test Performed | Checked: YES/NO |
| --- | --- | --- |
| 6. | PF/ Relay Contacts: Check for the relay change over for the set points. | Checked: YES |
| 7. | Display Annunciation: (As per the software Details) | Checked: YES |
| 8. | Communication: RS 232 / RS 485 / BCD / Other Communication Protocols (If any specify) The RS232 / RS485 to be checked with interfacing to PC | Checked: YES |
| 9. | Any other parameters (If Any) | N/A |
| 10. | Rating of the fuse used in the system: | 2A Fast Blow |

REMARKS (If ANY): - Calibration completed within acceptable industrial limits.

|  | LC 1 | LC 2 | LC 3 | LC 4 | LC 5 | LC 6 | LC 7 | LC 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SPAN | 10.00 | 10.00 | 10.00 | 10.00 | 10.00 | 10.00 | 10.00 | 10.00 |
| TARE | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 |
| EXC V | 10.0V | 10.0V | 10.0V | 10.0V | 10.0V | 10.0V | 10.0V | 10.0V |

Instrument Used: Fluke 87V Digital Multimeter & 50kg Calibrated Class F Weights

Tested by: A	                                                       		 Approved by: HOD-PDN

Date: 15-Sep-2026					                  		(HOD - PDN)