ROUTINE TEST RECORD FOR BELT SCALE

JOB No: JOB-2026-BS-001           					CUSTOMER NAME: ACME Pvt Ltd

CAPACITY : 500 TPH			             		CONVEYOR NO : CV-101

| BELT SCALE | LOAD CELL | SPEED SENSOR | JUNCTION BOX |
| --- | --- | --- | --- |
| Model No: BSW-4000 Serial No: SN-BSW-001 S/W Version: V2.5 REMOTE: Model No: RMT-100 Serial No: SN-RMT-001 | Model No: CAS-5000 Serial No: 1) LC-1001 5) LC-1005 2) LC-1002 6) LC-1006 3) LC-1003 7) LC-1007 4) LC-1004 8) LC-1008 | Model No: SPD-200 Serial No: SN-SPD-001 | Model No: JB-100 Serial No: SN-JB-001 Model No: Serial No: Model No: Serial No: |
|  |  | ANGLE SENSOR |  |
|  |  | Model No: Serial No: |  |

Check for short between Line & Neutral, Neutral & Earth, Line & Earth : Checked and found ok

| SL. No. | Test Performed | Specified | Actual |
| --- | --- | --- | --- |
| 1. | Input Supply Voltage (+/-10%) | 230V ±10% | 231V |
| 2. | Derived DC Voltages (+/-5%) | +5V, ±12V | +5.1V, ±12.2V |
| 3. | Display and Keypad Functionality | OK | OK |
| 4. | Belt Load as specified in G.A. or MPL | 190.48 | 190.50 Kg/m |
| 5. | Belt Speed as specified in G.A. or MPL | 1.75 | 1.76 m/s |
| 6. | Number of PF Contacts | OK | OK |
| 7. | Communication Output (Specify) | OK | OK |
| 8. | Analog Output (Specify number) | 2 x 4-20mA | 2 x 4-20mA |
| 9. | Wiring, TB and Component Layout | OK | OK |

Resolution:            L: 0.01          S: 0.01          r: 0.1          t: 0.1

Measurement to be done from 0 to 110% of the Belt Scale capacity.

| Load on Load cell Kg/Mv | Belt load Kg/m | Speed in m/s | Rate in tph | Totalizer in Tones (6 min) | O/p current mA |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | A01 | A02 |
| 0 | 0 | 0.00 | 0 | 0 | 4.0 |  |
| 50 | 25.0 | 1.75 | 157 | 15.7 | 8.0 |  |
| 100 | 50.0 | 1.75 | 315 | 31.5 | 12.0 |  |
| 150 | 75.0 | 1.75 | 472 | 47.2 | 16.0 |  |
| 200 | 100.0 | 1.75 | 630 | 63.0 | 20.0 |  |

Instrument Used: FLUKE 87 Series V

Tested by: John Doe                                				                           Approved by: Jane Smith

Date: 20/09/2026                            					                           (HOD - PDN)