TEST REPORT FOR ACCELEROMETER: MODEL NO PG109M0/PG114M0

Job No:

Customer Name:

Serial No:

Test 1. Axial Sensitivity

Linearity test

Frequency: 50Hz

Gain of the standard charge amplifier (G1) = 1.022mV/pC

| Acceleration from Std A (g) | Charge amplifier O/p of Test accelerometer V1 mV (Pk-Pk) | Charge O/p of Test accelerometer V2=V1G1 (pC) | Sensitivity of Test accelerometer S=V2A (pC/g) |
| --- | --- | --- | --- |
| 1 |  | {{spec_1}} | {{act_1}} |
| 2 |  | {{spec_2}} | {{act_2}} |
| 3 |  | {{spec_3}} | {{act_3}} |
| 4 |  | {{spec_4}} | {{act_4}} |
| 5 |  | {{spec_5}} | {{act_5}} |
| 6 |  | {{spec_6}} | {{act_6}} |
| 7 |  | {{spec_7}} | {{act_7}} |
| 8 |  | {{spec_8}} | {{act_8}} |
| 9 |  | {{spec_9}} | {{act_9}} |
| 10 |  | {{spec_10}} | {{act_10}} |

Mean:                pC/g

Frequency test

Gain of the standard charge amplifier(G1)= 1.022mv/pC

| Frequency (Hz) | Acceleration from Standard A (g) | Charge Amplifier output of Test Accelerometer V1 mV(Pk-Pk) | Charge output of test accelerometer V2=V1G1 (pC) | Sensitivity of test accelerometer S=V2A (pC/g) |
| --- | --- | --- | --- | --- |
| 50 | 1 | {{spec_1}} | {{act_1}} |  |
| 100 | 1 | {{spec_2}} | {{act_2}} |  |
| 200 | 1 | {{spec_3}} | {{act_3}} |  |
| 300 | 1 | {{spec_4}} | {{act_4}} |  |
| 400 | 1 | {{spec_5}} | {{act_5}} |  |
| 50 | 10 | {{spec_6}} | {{act_6}} |  |
| 100 | 10 | {{spec_7}} | {{act_7}} |  |
| 200 | 10 | {{spec_8}} | {{act_8}} |  |
| 300 | 10 | {{spec_9}} | {{act_9}} |  |
| 400 | 10 | {{spec_10}} | {{act_10}} |  |

Mean:             pC/g

Test 2. Post Sealing Axial Sensitivity

Linearity test

Frequency: 50Hz

Gain of the standard charge amplifier (G1) = 1.022mV/pC

| Acceleration from Std A (g) | Charge amplifier O/p of test accelerometer V1 mv (pk-pk) | Charge O/p of test accelerometer V2=V1G1 (pC) | Sensitivity of test accelerometer S=V2A (pC/g) |
| --- | --- | --- | --- |
| 1 |  | {{spec_1}} | {{act_1}} |
| 2 |  | {{spec_2}} | {{act_2}} |
| 3 |  | {{spec_3}} | {{act_3}} |
| 4 |  | {{spec_4}} | {{act_4}} |
| 5 |  | {{spec_5}} | {{act_5}} |
| 6 |  | {{spec_6}} | {{act_6}} |
| 7 |  | {{spec_7}} | {{act_7}} |
| 8 |  | {{spec_8}} | {{act_8}} |
| 9 |  | {{spec_9}} | {{act_9}} |
| 10 |  | {{spec_10}} | {{act_10}} |

Mean:         pC/g

| A) CHARGE SENSITIVITY | : |
| --- | --- |
| B) CAPACITANCE | : |
| C) INSULATION RESISTANCE | : |

| TESTED BY | : | APPROVED BY: |
| --- | --- | --- |
| DATE | {{cal_in_1}} | {{cal_out_1}} |