TEST RECORD FOR SIGNAL CONDITIONER

JOB ORDER No.	:	JOB-SC-007

CUSTOMER: Automation Systems Ltd

SIGNAL CONDITIONER					LOAD CELL

MODEL NO.: SC-4020 Dual Channel						MODEL NO.: LC-350-S

SERIAL NO.: SN-SC-8899						SERIAL NO.: SN-LC-1234

Linearity Test

| SL.NO. | LOAD ON LOAD CELL Kg/T OR mV | OUTPUT IN CURRENT / VOLTAGE |
| --- | --- | --- |
|  | 0.0 mV | 4.00 mA |
|  | 5.0 mV | 7.20 mA |
|  | 10.0 mV | 10.40 mA |
|  | 15.0 mV | 13.60 mA |
|  | 20.0 mV | 16.80 mA |
|  | 25.0 mV | 20.00 mA |

Check the POWER ON Annunciation: POWER ON Annunciation OK

Percentage of error: 0.02 %

Multimeter Used: Fluke 87V (Sl No: FLK-9901)

Tested by: Electronics Tech						Approved by: HOD-PDN

Date: 2026-08-22						(HOD - PDN)