TEST RECORD FOR TRIP SAFE

JOB ORDER No.	:	JOB-TS-008

CUSTOMER: Refinery Safety Unit

TRIPSAFE					LOAD CELL:	LC-S-BEAM-500

MODEL NO.: Trip Safe TS-300						MODEL NO.: LC-S-BEAM-500

SERIAL NO.: SN-TS-1234						SERIAL NO.: SN-LC-5501

Linearity Test:

| 1 | LOAD ON LOAD CELL |  | 0.00 mV | LED 1 (Green) |
| --- | --- | --- | --- | --- |
|  | 0 Kg | 0.0 mV |  |  |
| 2 | 200 Kg | 4.0 mV | 4.01 mV | LED 2 (Green) |
| 3 | 400 Kg | 8.0 mV | 7.99 mV | LED 4 (Yellow) |
| 4 | 600 Kg | 12.0 mV | 12.02 mV | LED 6 (Yellow) |
| 5 | 800 Kg | 16.0 mV | 15.98 mV | LED 8 (Red Trip) |
| 6 | 1000 Kg | 20.0 mV | 20.01 mV | LED 10 (Full Trip) |
| 6 |  |  |  |  |

Percentage of error for the Trip Level: 0.01 %

PF Contact for Healthy / Over Load: Healthy Relay Contacts OK

Lamp indication / LED annunciation as applicable: LED Bar Graph Annunciation OK

Instrument Used: Fluke 87V & Load Cell Calibrator

NOTE: Ensure that the Jumper/Shorting as per the Required Underload/Overload

Tested by: Safety Officer						Approved by: HOD-PDN

Date: 2026-08-22						(HOD - PDN)