Test Report for Miscellaneous items: Production Systems

Job No: JOB-MISC-9901 Model No: As Detailed Below Customer: JSW Steel Works Sl. No: As Detailed Below Item Description: As Detailed Below Qty: 7

Test Readings and Result: A. LOAD CELL RESISTANCE & INITIAL UNBALANCE TEST READINGS: (Model: LC-350-S-TYPE) Sl. No; Serial Number; Input Res (Ω); Output Res (Ω); Initial Unbal (mV/V); Status. 1; SN-LC-101; 350.2; 350.5; 0.02; PASSED. 2; SN-LC-102; 350.1; 350.4; 0.01; PASSED. 3; SN-LC-103; 350.4; 350.6; -0.03; PASSED. 4; SN-LC-104; 350.0; 350.3; 0.02; PASSED. Result: The above mentioned load cells were checked for input resistance, output resistance, and initial unbalance using Fluke 87V Digital Multimeter and were found OK. B. JUNCTION BOX INSPECTION: Model No: JB-04-SS-IP67 ¦ Type: 4-in 1-out Junction Box ¦ Qty: 1 ¦ Sl No: SN-JB-8821 Result: The above mentioned junction boxes were checked for cross-continuity, continuity, and output functionality and were found OK. C. SPEED SENSOR / TACHO INSPECTION: Model No: BCTWDS-SPD-PRO ¦ Indicator Model: LH32220 ¦ Qty: 1 ¦ Sl No: SN-SPD-1029 Result: The above mentioned speed sensor(s) were checked for speed with indicator model LH32220 and were found OK. D. REMOTE DISPLAY / INDICATOR INSPECTION: Type: 100mm Remote Display ¦ Model No: RD-100-LARGE ¦ Master Indicator: LS 32200 ¦ Sl No: SN-RD-901 Result: The above mentioned 100mm Remote Display unit(s) were checked for communication with indicator model LS 32200 and were found OK. E. ADDITIONAL INSPECTION & TEST FINDINGS:

Tested by: System Technician             		Approved by: Quality Manager

Date: 2026-08-22                         		(HOD-PDN)