Test Report for Miscellaneous items: Production Systems

Job No: JOB-MISC-9920 Model No: As Detailed Below Customer: Tata Steel Industrial Terminal Sl. No: As Detailed Below Item Description: As Detailed Below Qty: 6

Test Readings and Result: A. LOAD CELL RESISTANCE & INITIAL UNBALANCE TEST READINGS: Sl. No; Load Cell Model; Serial Number; Input (Ω); Output (Ω); Unbal (mV/V); Status. 1; LC-350-COMPRESSION; SN-LC-350-01; 350.2; 350.5; 0.02; PASSED. 2; LC-350-COMPRESSION; SN-LC-350-02; 350.1; 350.4; 0.01; PASSED. 3; LC-500-TENSION; SN-LC-500-01; 350.4; 350.6; -0.03; PASSED. 4; LC-500-TENSION; SN-LC-500-02; 350.0; 350.3; 0.02; PASSED. Result: The above mentioned load cells were checked for input resistance, output resistance, and initial unbalance using Fluke 87V Digital Multimeter and were found OK. B. JUNCTION BOX INSPECTION: Sl. No; Junction Box Type; Model Number; Serial Number(s); Qty; Status. 1; 4-in 1-out Junction Box; JB-04-SS-IP67; SN-JB-401; 1; PASSED. 2; 6-in 1-out Junction Box; JB-06-SS-IP67; SN-JB-601; 1; PASSED. Result: The above mentioned junction boxes were checked for cross-continuity, continuity, and output functionality and were found OK. C. SPEED SENSOR / TACHO INSPECTION: Sl. No; Equipment Description; Model Number; Indicator Model; Serial Number; Status. 1; Speed Sensor / Tacho; BCTWDS-SPD-PRO; LH32220; SN-SPD-1029; PASSED. Result: The above mentioned speed sensor(s) were checked for speed with indicator model LH32220 and were found OK. D. REMOTE DISPLAY / INDICATOR INSPECTION: Sl. No; Display Type / Size; Model Number; Master Indicator; Serial Number; Status. 1; 100mm Remote Display; RD-100-LARGE; LS 32200; SN-RD-901; PASSED. Result: The above mentioned 100mm Remote Display unit(s) were checked for communication with indicator model LS 32200 and were found OK. E. ADDITIONAL INSPECTION & TEST FINDINGS:

Tested by: System Technician             		Approved by: Quality HOD

Date: 2026-08-22                         		(HOD-PDN)